The EEG sensor is an electroencephalograph sensor, or pre-amplifier, for measuring the brain’s electrical activity from the surface of the scalp. It amplifies the small electrical voltages that are generated by brain cells (neurons) when they fire. Similarly to muscle fibers, neurons of different locations can fire at different rates. The frequencies most commonly looked at, for EEG, are between 1 and 40 Hz. The EEG sensor records a “raw” EEG signal, which is the constantly varying difference of potential between the positive and negative electrode, and the software usually processes that signal by applying a variety of digital filters to the recorded signal, in order to extract frequency-domain information. EEG signals are usually measured in microvolts (μV).

**SENSOR PLACEMENT**

International 10-20 system for placement [1] is recommended. Skin preparation and use of adhesive, conductive gel is also most often recommended for best results.

**Skin preparation:** For any recording where signal quality matters, good skin preparation is important to get a clean signal and avoid artifacts. Before applying the EEG electrodes, use a skin preparation cream, such as NuPrep. If necessary, shaving excess body hair may be required (see instructions below).

**Electrodes Selection:** the DIN interface cable (T8740) allows any EEG electrode with 1.5 mm standard DIN connector to be connected to the EEG sensor. For single channel recordings, three (3) electrodes will be needed (one active, one reference and one ground). When recording multiple channels simultaneously, jumper cables such as SA9315-2 and SA9315-4 can be used to allow a single electrode to be used as a common reference and/or common ground. This reduces the number of electrodes needed. Most often silver-silver/chloride electrodes are used for best performance although many types of electrodes will work adequately (gold or tin, for example).

NuPrepGood Electrode Contact: Conductive gel is required for optimal electrode-skin contact. Make sure the electrodes are placed firmly on the skin and that there is good contact between skin and electrode (see instructions below).


**10-20 electrode placement system**

The electroencephalogram (EEG) is a recording of the electrical activity of the brain, most often captured at the surface of the scalp. The ten-twenty (10-20) electrode system of the International Federation [1] is the standard for electrode placement. It is used to place surface EEG electrodes in a repeatable way independent of inter-subject anatomical variability.

It is called 10-20 because of the way distances between electrode sites are computed. The distances between certain anatomical landmarks are segmented at increments of 10% and 20% of their value, and electrodes are placed at these points.

The sites are named as follows:


* A letter is used to indicate over which area of the brain the site is located. These are:
  + F: frontal lobe
  + FP: front polar area
  + C: central sulcus
  + T: temporal lobe
  + P: parietal lobe
  + O: occipital lobe
  + Other miscellaneous labels are used for the ears (A) and other reference sites (for example, M for mastoid process, G for ground, etc…)
* The letter Z is used to indicate the central line along the interhemispheric fissure.
* A number is used to indicate the position in reference to the central line (Z).
  + Odd numbers are on the left
  + Even numbers are on the right
  + The number value increases the further away the site is from the central line.



or example, along the line joining sites A1 and A2: to the right of A1, at 10% of the overall A1-A2 distance is electrode site T3. This is followed by C3 (20% further), Cz (20% further), C4 (20% further), and T4 (20% further). Site T4 should fall 10% to the left of A2.

**EEG electrodes**

For one channel EEG recordings, only one T8750 kit is required.



**Monopolar Configuration**

Connect the blue active electrode to the blue DIN cable input.

Connect the yellow ear clip to the yellow DIN cable input, and the black earclip to the black DIN cable input.

The other end of the DIN cable plugs into the input of the EEG sensor.

**Note:** For bipolar configuration, the yellow cup electrode is used in place of the yellow earclip.

**EEG electrode placement technique**

Proper site preparation and electrode placement are essential for clean, trouble-free recording of the EEG signal. The following steps describe how to achieve a secure connection in an efficient, repeatable way. They should be followed closely every time the EEG signal is recorded. This may seem daunting at first, but the technique is easily learned and the user will quickly become accustomed to it and will certainly appreciate the benefits.


1. Locate the required electrode sites.

(In this example the site Cz will be used.)

* Locate the Inion. On the posterior base of the skull, where the spine meets the skull, there is a small protrusion called the “inion”. It can be found by running the finger up the spine towards the skull.

* There is a small crevice between the spine and skull in which vertebrae can no longer be felt. Just above that area, the inion of the skull can be found.

* Locate the nasion. The depression on the bridge of the nose, just below the brow and directly between both eyes, is called the “nasion”.


* The line between these two points runs along the interhemispheric fissure of the brain (the space that separates the left and right sides of the brain, i.e. right down the middle).


Mark a spot located at 50% of the distance between the nasion and the inion.

* Locate the mandibular notch.

Anatomy revisited: To locate the mandibular notch, place a finger against the tragus, just anterior to the ear, and ask the subject to open their mouth. Your finger should find its way into a cavity, resting superior to the manibular notch.

Find this spot on both sides of the head.

* The line that connects the left and right mandibular arches runs along the central sulcus of the brain (the space that separates the frontal and parietal lobes).
* Mark a spot located at 50% of the distance between the left and right mandibular notches. It should intersect directly with the mark made between the front and back of the head.
* The intersection of these two lines is electrode location Cz.



**NOTE:** Electrode location CZ can always be approximated instead of explicitly measured following the procedure described above. It is the top middle of the skull. Imagine if the subject were to hang perfectly straight from a string tied to the top of their heads, this is where the string would have to be fixed. Still, to ensure repeatable measurements, especially when comparing across sessions, it is recommended that the electrode locations be measured as described.

**2. Prepare the scalp.**

The scalp must be prepared before an electrode can be fixed. This involves slightly abrading the skin to remove dead skin, sweat and other contaminants to the EEG signal.

* Scoop up a small quantity of NuPrep™ skin prepping gel on a cotton swab or tissue.

* With the thumb and index finger of one hand, separate the hair around the electrode site that was previously found and marked.


* Run the gel in the direction of the natural line formed along the scalp by the split hair. Some light force must be used, enough to redden the scalp slightly, but not enough to break the skin.

* 

Wipe away the excess prepping gel with a dry, lint-free cloth. Care should be taken to keep the hair parted and to keep track of the site after wiping clean.

**Helpful trick:** A small amount of Ten20™ Conductive paste should be smeared on the newly prepared site. This has the combined effect of keeping track of the site, keeping the hair neatly out of the way, and acting as a landing pad for the electrode once it is ready to fix.

**3. Prepare the electrodes**



* Fill the electrode cup with Ten20™ conductive paste so that no air bubbles exist in the cup.

* Add more Ten20 paste onto the cup electrode, just enough to form a ball on the cup, not so much that it spills over the edge. Shown is the ideal amount of paste.

* 

Place the cup face down on the landing pad previously prepared. Gently push the electrode down to fix it to the scalp. A little bit of paste should run out along the edge of the cup to form a thin ring around it.

Helpful trick: Place the electrode so that the direction of the cable does not place undue stress on the cup (so that it gets pulled, lifted or twisted off). The cable should hang naturally and towards the plastic clip (as shown). Leave enough slack in the cable to allow for comfortable head movement as well.

**4. Reference sites.**

* Repeat the above preparation and placement steps on the ear lobes. Do not put too much paste on the ear clip electrode, but ensure that the gold disc is completely covered.



Leave enough slack in the cable to allow the subject to turn their head easily, but not too much that it can get caught. Think about the angle of the cable from the ear clip to the neck clip, so that no extra tension is placed on the ear clip.

Viewed from behind, the final configuration should look like this:

Note the position of the clip, the direction of the cables and the slack left to provide mobility.

**References**

[1] H.H. Jasper, “The ten-twenty electrode system of the International Federation”, Electroencephalography and Clinical Neurophysiology, vol. 10, pp. 370-375, 1958.

**TECHNICAL SPECIFICATIONS**

* Size : 37mm x 37mm x 12mm (1.45" x 1.45" x 0.45")
* Weight : 25g (0.5 oz)
* Input Impedance : 10GΩ in parallel with 10pF
* Signal Input Range : 0 – 200µV
* Sensitivity : < 0.1 µVRMS
* CMRR : > 130 dB
* Channel Bandwidth : 2 Hz-1 kHz
* Accuracy : ±0.3 µVRMS, ±5% of reading @10°C to 40°C

**ELECTRICAL COMPATIBILITY**

The EEG sensor is designed to coexist with other Thought Technology bio potential sensors such as T9306M(or T9307M) EKG sensor, T9303M MyoScan sensor, T7680 EEG-Z3 sensor, or SA9309M Skin Conductance sensor.

To ensure correct EEG sensor operation, if sensors from another manufacturer are in the same electrical circuit and connected to the same subject, their electrodes must function at a voltage within the specified operating bias range, 1.0 to 3.0 volts above sensor ground. To check whether another sensor is interfering with the EEG sensor operation, connect and disconnect the other sensor from the subject, and note whether this causes a change in the EEG sensor signal level, or whether connection of the other sensor appears to cause any signal artifacts in the EEG signal.

**INTERFACING WITH 3RD PARTY DATA ACQUISITION SYSTEMS**

**Recommended Connectivity for Electrical Safety**


To ensure electrical safety in the user setup, Thought Technology recommends the use of TT Sensor Isolator ST9405AM when interfacing patient connected sensor(s) to line powered equipment(s) or devices.

The TT Sensor Isolator ST9405AM is an interface device providing medical grade electrical isolation between the patient connected sensors and the acquisition system. It provides the equivalent of Two Means of Patient Protection under IEC 60601-1, and supplies battery power to the sensors. Using this device ensures Thought Technology sensors are safely interfaced to the analog inputs of line-powered systems such as computers with DAQ cards.

**Note that this device isolates only between sensors and the DAQ interface, not between different sensor channels.**

The TT Sensor Isolator can interface up to 4 sensors to a DAQ card. TT Sensor Isolator can be connected to the DAQ card in two ways:

* via two stereo jacks, or
* via a DB-15 connector; a BNC interface cable (SA9409BNC) or a pigtail cable (SA9409PGT) can be provided with the unit.



For more detailed information on the Sensor Isolator 4∞, consult the Thought Technology Science Division website or contact the sales department or your distributor.  
Division website or contact the sales department or your distributor.


**Direct Connectivity for Electrically Isolated Systems**

The following notes are provided for qualified users to directly interface Thought Technology sensors with external systems.

**WARNING: If the sensor is interfaced to non-Thought Technology devices without the use of a TT Sensor Isolator SE9405AM, an elevated risk of electrical shock may be present. In particular, if a patient-connected sensor is connected to any line powered device(s) or equipment(s), it will be the responsibility of the qualified user to ensure the electrical safety in the setup and to ensure that the device or equipment provides sufficient isolation.**

To interface with a sensor, a single sensor cable may be cut in half. Both sides can then be used to make custom interfacing cables by stripping the outer insulation of each required conductor. The sensor cable contains 4 color coded conductors. The table below shows the color coding and pin connector assignment.

|  |  |  |  |
| --- | --- | --- | --- |
| **Pin** | **Color Code** | **Function** | **Note** |
| 1 | metal (shield) | ground | Signal and power ground, connection required. |
| 2 | yellow | auxiliary (sensor ID) | No connection required. |
| 3 | green | signal | Sensor output signal |
| 4 | red | sensor power | Supply voltage, +7.26V referenced to ground. Note: sensor performance may be sensitive to supply voltage. |



**Notes:**

The nominal supply voltage for this sensor is 7.26V. The sensor can safely be used with a supply voltage of up to 9V. However, as the sensor is calibrated with a 7.26V supply voltage level, changes in gain and offset may be expected when operating at a different supply voltage level, changes in gain and offset may be expected when operating at a different supply voltage.

**General Specifications and Recommendations**

* Recommended minimum 16-bit ADC
* Input range:
  + Connected using the product SE9405AM Sensor Isolator: 0-5V (unipolar) or ±5V
  + Direct connect with AC coupled output and DC bias resistor: ±1.25V or ±2.5V

**Simplified Transfer Function**

equationOutput voltage (input to DAQ) from input voltage (raw EEG signal)

equationOutput voltage (input to DAQ) from input voltage (raw EEG signal)

**Notes:**

* The highpass effect of the sensor’s AC coupling is not shown in the transfer function. The function is accurate for frequency components within the specified bandwidth, and should adequately represent the scaling of an EEG signal.
* The 2.8 volt term is present if the TT Sensor Isolator is used. It is absent, if the sensor is directly connected to a data acquisition card, and if a resistor is connected across its output (see section below).

**SPECIFICATION SUMMARIES OF SUPPORTED ACCESSORIES / HARDWARE**

Table below lists Thought Technology accessories for the EEG Sensor.

|  |  |
| --- | --- |
| **Accessory** | **Product Number** |
| Monopolar / Bipolar kit | Monopolar / Bipolar kit (T8750)  This kit allows users to make either monopolar or bipolar single channel measurements.  For customers who already have the EEG DIN cable, this kit is also available as T8751 without the EEG DIN cable at a lower cost.   |  |  |  | | --- | --- | --- | | Qt. | PN# | Description | | 1 | SA8740 | EEG Extender DIN cable | | 1 | SA9323B | TT-EEG gold cup cable – blue | | 1 | SA9323Y | TT-EEG gold cup cable – yellow | | 1 | SA9321 | TT-EEG gold ear clip – black | | 1 | SA9321Y | TT-EEG gold ear clip – yellow | |
| TT-EEG Gold Cup Cable | TT-EEG Gold Cup Cable (SA9323)  Available in black (SA9323), yellow (SA9323Y), and blue (SA9323B) |
| TT-EEG Gold Ear Clip | TT-EEG Gold Ear Clip (SA9321)  Available in black (SA9321) and yellow (SA9321Y) |
| TT-EEG Gold Ear Clip | EEG DIN Extender Cable (T8740)  For connection of three standard DIN terminated. |
| 1M2F Y-Connector (SA9315) | 1M2F Y-Connector (SA9315)  Connects a single electrode to two inputs. Allows use of the same electrode in two channels for common ground or reference. |
| 1M4F Y-Connector (SA9315-4) | 1M4F Y-Connector (SA9315-4)  Connects a single electrode to four inputs. Allows use of the same electrode in up to four channels for common ground or reference. |
| 2M1F Y-Connector (SA9319) | 2M1F Y-Connector (SA9319)  Connects two electrodes to a single input. Used to join two ear clip cables together to form a linked ear reference. |

Table of Contents

* [**Product Overview**](http://www.thoughttechnology.com/sciencedivision/pages/products/eegflex.html)
* [**Technical Specifications**](http://www.thoughttechnology.com/sciencedivision/pages/products/eegflex.html)
* [**Interfacing with 3rd Party Data Systems**](http://www.thoughttechnology.com/sciencedivision/pages/products/eegflex.html)
* [**Summary of Supported Hardware**](http://www.thoughttechnology.com/sciencedivision/pages/products/eegflex.html)

Suggested References